admin\manual\preliminary-vessel

SUPERPRESSURE

#### PRODUCTS

PRELIMINARY – TEST

##### REACTION VESSELS

CATALOG NO. 41-19230,

41-19250,41-19270

NEWPORT SCIENTIFIC, INC.

FORMERLY - AMINCO

8246-E SANDY COURT

JESSUP, MD 20794

(301) 498-6700   FAX (301) 490-2313

TABLE OF CONTENTS

INSTALLATION									3

OPERATION									3

ASSEMBLY PROCEDURE							3

DISASSEMBLY PROCEDURE							4

MAINTENANCE									4

PART LIST									6

FIGURES

1	PRELIMINARY-TEST REACTION VESSEL			2

2	REACTION VESSEL CROSS-SECTIONAL VIEW   	6

1

These vessels are used for making simple and convenient preliminary tests of chemical reactions.  The body of a reaction  vessel is bored from solid bar stock.  Two flats at the bottom of the reaction vessel permit it to be held with a wrench or in a vise to facilitate removal of the head.  The closure is a screw-plug which is seated on a flat metallic gasket made of copper or stainless steel.  A threaded opening in the head and another in the bottom of the reaction vessel are provided for connecting to 1/4-inch o.d. NSI tubing.  The openings are supplied with nuts and sleeves for connecting to tubing.  A 3/8-inch or 9/16-inch tubing connection can be furnished on special order.

The following apply to all models:

-  Approximate weight:  8 lbs

-  Outside diameter:  2 inches

-  Inside diameter:  1 inch

-  Inside depth:  8 inches

FIGURE 1. PRELIMINARY-TEST

REACTION VESSEL		-  Approx. Volume of reaction vessel:

103ml

The characteristics of each reaction vessel are as follows:

2

| Cat No. | Mode of Construction | Maximum Working Pressure (lb/sq in.) |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  |  | ##### 100F | ##### 500F | ##### 650F | ##### 800F |
| 41-19230 41-19250 41-19270 | A.I.S.I. 316 A.I.S.I. 410 Chrome-Vanadium Steel | 11,250 12,750 20,250 | 10,300 11,200 16,880 | 10,250 10,750 12,800 | 10,080 9,800 6,780 |

###### INSTALLATION

Unpack reaction  vessel and inspect for damage.  Follow the guide line given in the “operation” paragraph to install and use the reaction vessel for your application.

###### OPERATION

The reaction vessel can handle solid, liquid, or gas samples in any combination.  The operating procedure will vary according to the specific application.  In general:

-  If using a sealed system, follow the guideline in the “Assembly Procedure” and introduce samples into the reaction vessel according to your application.

-  If you must open the reaction vessel to introduce the sample, follow the “Disassembly Procedure”, introduce the sample, and then follow “Assembly Procedure” to seal the reaction vessel.

ASSEMBLY PROCEDURE

1.  Tighten head on reaction vessel to a torque of

approximately 225 foot-pounds.

2.  Connect ¼-inch O.D. tubing to the threaded opening in

the head and the bottom of the body secure with gland

nuts.

3.  Tighten these gland nuts to a torque of approximately

15 foot-pounds.

3

4.  Check to ensure that all connections are secure.

The reaction vessel is now assembled.

###### DISASSEMBLY PROCEDURE

1.  RELEASE PRESSURE FROM THE REACTION VESSEL.

## WARNING

BEFORE DISCONNECTING TUBING FROM THE REACTION VESSEL,

ENSURE THAT THE GAS PRESSURE IS RELEASED.  IF CONTENTS

OF REACTION VESSEL ARE HAZARDOUS OR IRRITATING, WEAR

PROTECTIVE GLASSES AND GLOVES.

2.  Loosen gland nuts and remove tubing connections.

3.  Using a syringe with a large bore needle, inject

solvent into the reagent vessel.

4.  Plug openings and vigorously shake reaction vessel.

5.  Remove plugs, rinse and drain.  Repeat from step

3, as necessary.

6.  Remove any residue from exterior surface.

7.  Secure bottom of reaction vessel in a vise and

unscrew head.

8.  If required, reclean vessel and head.

###### MAINTENANCE

Maintenance consists of keeping the reaction vessel free from contaminants, lubricating the threads between the body and head, and replacing the head gasket, whenever necessary.  The procedure follows:

1.  Disassemble reaction vessel as described.

4

2.  Clean using hot tap water and applicable solvents

and /or detergent.  A bristle brush may be used to

remove stubborn residues.

3.  Rinse with a volatile low-residue solvent or

distilled or deionized water, if water cleaned.

4.  Dry with a lint free wipe and/or rinse with acetone

or similar solvent.  Air dry to insure removal of all

liquids.

5.  If stubborn residues remain, use a wood or plastic

scraper to remove them and then reclean surfaces

and dry.

6.  Make certain that the head gasket is free of nicks

or scratches – particularly radial ones.  This gasket

Fits into recess machined into the underside of the

head.  Stainless-steel vessels have stainless steel

gaskets which may not be distinguishable from the head.

copper gaskets are easily distinguished.  If necessary,

replace gasket as follows:

A.  Spring the gasket from the groove by striking

it sharply with a cape chisel around the inner

and outer circumferences.  Do not chisel down

on the face of the gasket or damage to the head

will result.

B.  Drop a new, clean gasket into the groove in

the head.  A slight clearance should exist

between the gasket and groove.  After several

installations and removals of the head, the

gasket will expand into the recess and a ridge

will appear on the gasket.

7.   Lubricate the threads on the body and head with

lubriplate 130A (62147002300).  Every precaution

should be taken to  prevent this lubrication from

entering the internal body of the reaction vessel.

8.  Assemble the reaction vessel as described in the

assembly procedure.

5

PARTS LIST (See Figure 2)

| REF NO. | ## DESCRIPTION | Part No. |  |  |
| --- | --- | --- | --- | --- |
|  |  | 41-19230 | 41-19250 | 41-19270 |
| 1 2 3 4 5 | GLAND NUT (2REQ.) INNER SLEEVE(2REQ.) BODY GASKET* HEAD | 45-11313 45-11316 53063000200 P1604014900(S.S) 5306300300 | 45-11313 45-11316 53063000203 P1604014900(S.S) 53063000303 | 45-11313 45-11316 53063000201 P1604014800(COPPER) 53063000301 |
| * The stainless steel and copper gaskets can be used interchangeably. |  |  |  |  |

FIGURE 2. REACTION VESSEL CROSS-SECTIONAL VIEW

6